MAT-DOC

Mounting Assistive Technology Documentation

Version 2.2

[Enter Name in Properties]

# Section 1 – Pre-Assessment Information

| SUMMARY INFORMATION |  |  |
| --- | --- | --- |
| Client Name | Form Completion Date | Form Completed By |
| [Enter Name in Properties] |  |  |
| Organisation | Email Address | Phone Number |
| [Enter service name in properties] |  |  |

| CLIENT INFORMATION |  |  |  |
| --- | --- | --- | --- |
| Client’s weight | kg | Is a change in weight expected? |  Decrease  No change expected  Increase |
|  |  Estimated  Measured (Date: ) |  |  |
| Transfer method |  Hoist  Standing  Side (Left / Right)  Other | Transport method |  None  Taxi  Wheelchair Adapted Vehicle  School Bus / Public Transport |
| Additional Comments and Observations |  |  |  |

| EQUIPMENT TO BE MOUNTED |  |  |  |
| --- | --- | --- | --- |
| Devices and Accessories | Access Method | Mounting Plate | Total Weight |
|  |  |  Daessy  REHAdapt  Universal  Other | kg |

| ENVIRONMENTAL CONSIDERATIONS Consider a typical week in the life of the device user, and where the device is likely to be used. |  |  |
| --- | --- | --- |
| Home | Where is time spent and use of the device required? |  |
|  | What wheelchairs/seating systems are used? |  |
|  | Will the device need to be accessed from bed? |  |
| School / Work | Where is time spent and use of the device required? |  |
|  | What wheelchairs/seating systems are used? |  |
|  | Will the device be accessed at any desks? |  |
| Outside/ Leisure | Where is time spent and use of the device required? |  |
|  | What is the terrain like? e.g. slopes/uneven |  |
|  | What wheelchairs/seating systems are used? |  |
| Other | Describe any other locations where the technology may be required? |  |
|  | Is there any other equipment that should be considered (lying support, standing support, or walker for example) |  |

| PHOTOS |  |
| --- | --- |
| Overall View | Device Held in Place By Hand |
| [Insert photo of overall view here] | [Show the device and access devices held in place] |
| Final Mounting Solution(s) |  |

# Section 2 – Wheelbase Information

|  | WHEELBASE 1 | WHEELBASE 2 | WHEELBASE 3 |
| --- | --- | --- | --- |
| If manual, how is the chair propelled? |  Self  Attendant |  Self  Attendant |  Self  Attendant |
| If powered, how is the chair controlled? |  Joystick  Switches |  Joystick  Switches |  Joystick  Switches |
| Joystick / Switch site |  |  |  |

|  | WHEELBASE 1 | WHEELBASE 2 | WHEELBASE 3 |
| --- | --- | --- | --- |
| Wheelbase Owner |  |  |  |
| Owner Contact Phone |  |  |  |
| Owner Asset ID |  |  |  |
| Is the chair due to be reviewed within the next six months? |  |  |  |

|  | WHEELBASE 1 | WHEELBASE 2 | WHEELBASE 3 |
| --- | --- | --- | --- |
| Manufacturer |  |  |  |
| Model |  |  |  |
| Serial number |  |  |  |
| Max user weight (kg) |  |  |  |

|  | WHEELBASE 1 | WHEELBASE 2 | WHEELBASE 3 |
| --- | --- | --- | --- |
| Any obstacles to mounting (e.g. joystick or tray)? |  |  |  |
| Tilt in Space? |  |  |  |
| Recline? |  |  |  |
| High-Low Base? |  |  |  |

# Section 3.1 – Wheelchair Mounting

| Have any issues been identified when considering… | Y / N | Notes |
| --- | --- | --- |
| …the types of terrain on which the chair will be used (indoors, outdoors, uneven, smooth)? |  Yes  No |  |
| …the desired position of equipment (considering user and their set-up / environment)? |  Yes  No |  |
| …the clients’ method of transfer, and whether this is likely to change in the future? |  Yes  No |  |
| …whether the device will be in a fixed position or will require adjustment of tilt or angle? |  Yes  No |  |
| …whether the device will need to fold or swing away to the side or behind the chair when not in use? |  Yes  No |  |
| …whether the position of the device will need to be altered for driving? |  Yes  No |  |
| …whether quick removal of the device from the mount or chair will be required? |  Yes  No |  |
| …whether the device will be moved, adjusted or removed by the user? |  Yes  No |  |
| …whether the mounting system is compatible with any existing systems? (other chairs, floor or desk mounts)? |  Yes  No |  |
| …whether the chair needs to be used at a desk or table? |  Yes  No |  |
| …whether carer activities (such as feeding) are performed from a particular side? |  Yes  No |  |
| …the rigidity of the host wheelchair / wheelbase? |  Yes  No |  |

DECISION TABLE: Wheelchair Mounting

|  | Wheelbase 1 | Wheelbase 2 | Wheelbase 3 |
| --- | --- | --- | --- |
| Mount Manufacturer and Type |  Daessy (Standard)  Daessy (M-Series)  Daessy (Lite Mount)  REHAdapt (M3D)  REHAdapt (L3D)  Manfrotto  Other  Bespoke / Hybrid  Side-Folding  Rear-Folding  Fixed |  Daessy (Standard)  Daessy (M-Series)  Daessy (Lite Mount)  REHAdapt (M3D)  REHAdapt (L3D)  Manfrotto  Other  Bespoke / Hybrid  Side-Folding  Rear-Folding  Fixed |  Daessy (Standard)  Daessy (M-Series)  Daessy (Lite Mount)  REHAdapt (M3D)  REHAdapt (L3D)  Manfrotto  Other  Bespoke / Hybrid  Side-Folding  Rear-Folding  Fixed |
| Mount Point (with reference to user) |  Left  Right  Central |  Left  Right  Central |  Left  Right  Central |
| Mount Component List |  |  |  |
| Mount Component Total Weight |  |  |  |
| Notes |  |  |  |

| If the wheelchair / wheelbase has any of the following features, have any issues been identified when considering… | Base 1 | Base 2 | Base 3 | Notes |
| --- | --- | --- | --- | --- |
| …tray? |  Yes  No  N/A |  Yes  No  N/A |  Yes  No  N/A |  |
| …tilt and / or recline functionality? |  Yes  No  N/A |  Yes  No  N/A |  Yes  No  N/A |  |
| …riser or high-low base? |  Yes  No  N/A |  Yes  No  N/A |  Yes  No  N/A |  |
| …footplates? |  Yes  No  N/A |  Yes  No  N/A |  Yes  No  N/A |  |
| …brakes? |  Yes  No  N/A |  Yes  No  N/A |  Yes  No  N/A |  |
| …folding seat or wheelbase? |  Yes  No  N/A |  Yes  No  N/A |  Yes  No  N/A |  |
| …tie-down points for transportation in a vehicle? |  Yes  No  N/A |  Yes  No  N/A |  Yes  No  N/A |  |

Based on decisions above, ensure available capacity is greater than 0 using table below:

| Wheelbase Ref (from Section 2) | Weight Limit (A) | Weight of AT Device and Any Other Equipment * (B) | Weight of Mount (C) | Weight of User (D) | Available Capacity (A – B – C – D) ** |
| --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |
| 2 |  |  |  |  |  |
| 3 |  |  |  |  |  |

* This should include the equipment being mounted and any other items required for medical purposes or personal care which are adding extra weight to the chair. Such items may be permanently attached or removable.

** Ensure that the available capacity is greater than zero. Where the available capacity is less than 10% of the chair’s total weight limit, proceed with caution and consult the wheelchair service.

| Advice given on capacity: |  |
| --- | --- |

WHEELCHAIR STABILITY SCREENING - FOR GUIDANCE ONLY
The guidance given by this screening tool is not based on empirical data but is instead based on the experience and professional judgement of the MAT-DOC Working Party. This screening tool should only be used as a guide and is not intended to replace a full risk assessment or risk management programme.

|  |  |  |  |  |  |  | Base 1 Score | Base 2 Score | Base 3 Score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Base Type | Powerchair |  | Manual |  | Hi-Lo chair |  |  |  |  |
|  | 1 |  | 3 |  | 5 |  |  |  |  |
| Location | Inside only |  | Used outside |  | Steep environments |  |  |  |  |
|  | 1 |  | 2 |  | 5 |  |  |  |  |
| User Movement | Moves within base footprint |  | Moves outside base footprint 1 |  | Moves significantly outside base footprint 2 |  |  |  |  |
|  | 1 |  | 3 |  | 5 |  |  |  |  |
| Kit Weight 3 | < 2kg |  | 2 – 5kg |  | > 5kg |  |  |  |  |
|  | 1 |  | 2 |  | 3 |  |  |  |  |
| Kit Position | Within footprint |  | < 20cm outside footprint |  | > 20cm outside footprint |  |  |  |  |
|  | 1 |  | 2 |  | 5 |  |  |  |  |
| Wheelchair Seating | Manufacturer’s standard seating only | Seat / pressure cushion in situ |  | Special seating |  | Special seating in atypical position 4 |  |  |  |
|  | 1 | 2 |  | 3 |  | 5 |  |  |  |
| TOTAL |  |  |  |  |  |  |  |  |  |

|  | 1 Centre of gravity remains inside base |  | 3 Weight of the assistive technology being mounted and the mount itself |  |  |
| --- | --- | --- | --- | --- | --- |
|  | 2 Centre of gravity of user moving outside of base |  | 4 If the special seat is mounting in a forward position or rearward position. |  |  |
| SCORE |  | ACTION |  | DECISION REASONING |  |
| <= 9 |  | Standard risk management required |  |  |  |
| 10 - 18 |  | More detailed risk management should be considered |  |  |  |
| 19 + |  | Detailed risk management must be completed, agreed and documented – consider full stability screening (Appendix 1) |  |  |  |

# Section 3.2 – Desk / Table Mount

| In considering the following, have any issues been identified? | Y / N | Notes |
| --- | --- | --- |
| …the size and stability of the surface on which the mount will be placed? |  Yes  No |  |
| …the desired position of AAC device (considering user and their set-up / environment)? |  Yes  No |  |
| …the possibility that the mount could be placed in an unsafe position or location? |  Yes  No |  |
| …whether the device will be in a fixed position or will require adjustment of tilt or angle? |  Yes  No |  |
| …whether the device will be moved or adjusted by the user? |  Yes  No |  |
| …whether the mounting system is compatible with any existing systems? (chair, floor or desk mounts)? |  Yes  No |  |
| …whether the mount impedes any other activities? |  Yes  No |  |
| …whether the mount impedes care activities (such as feeding)? |  Yes  No |  |

|  | Desk Mount |  |  |
| --- | --- | --- | --- |
| Mount Manufacturer and Type |  Daessy Folding Desk Mount  Daessy Lite Mount  REHAdapt Folding Desk Mount  REHAdapt L3D  Manfrotto  Other | Total Weight of Equipment |  |
| Notes |  |  |  |

# Section 3.3 – Floor Mount

| Have any issues been identified when considering… | Y / N | Notes |
| --- | --- | --- |
| …the environment where the floor mount will be used? |  Yes  No |  |
| …the manoeuvrability or any obstructions to the desired position of the mount? |  Yes  No |  |
| …where the floor mount will be stored when not in use? |  Yes  No |  |
| …the size and stability of the surface on which the mount will be placed? |  Yes  No |  |
| …the desired position of AAC device (considering user and their set-up / environment)? |  Yes  No |  |
| …the possibility that the mount could be placed in an unsafe position or location? |  Yes  No |  |
| …whether other people in the environment present an increased risk? |  Yes  No |  |
| …whether the device will be in a fixed position or will require adjustment of tilt or angle? |  Yes  No |  |
| …whether the device will be moved or adjusted by the user? |  Yes  No |  |
| …whether the mounting system is compatible with any existing systems? (chair, floor or desk mounts)? |  Yes  No |  |
| …whether the mount impedes any other activities? |  Yes  No |  |
| …whether the mount impedes the operation of other equipment (such as beds or armchairs)? |  Yes  No |  |

|  | Floor Mount |  |  |
| --- | --- | --- | --- |
| Mount Manufacturer and Type |  Daessy Rolling Mount  Daessy Mini- / Midi- Floor Mount  REHAdapt Classic Floor Mount  REHAdapt Pentalock  REHAdapt Mini  Other | Total Weight of Equipment |  |
| Notes |  |  |  |

# Section 4 – Hazard Identification and Risk Management

This section is designed to aid the identification of hazards which may occur due to the provision of a mounting system and to provide suggestions for risk management. Hazards should be identified and scored using the risk matrix below (or local equivalent, as required). This section uses the definitions outlined in BS EN ISO 14971 Medical Devices – Application of Risk Management to Medical Devices, which are as follows:

Hazard – a potential source of harm

Risk – the combination of the probability of occurrence of harm and the severity of that harm

Risk should be calculated pre- and post-management to demonstrate that the management procedures put in place have reduced the likelihood of an identified hazard occurring. The following formula is suggested for these calculations:

S – Severity

P – Probability

T – Total (T = S x P)

If using a bespoke or hybrid mount, a full technical file and risk management documentation should be kept, in accordance with local policies.

| SAMPLE 5 x 5 RISK MATRIX |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Probability | Severity |  |  |  |  |
|  | 1 (Negligible) | 2 (Minor) | 3 (Moderate) | 4 (Severe) | 5 (Significant) |
| 1 (Remote) | 1 | 2 | 3 | 4 | 5 |
| 2 (Unlikely) | 2 | 4 | 6 | 8 | 10 |
| 3 (Possible) | 3 | 6 | 9 | 12 | 15 |
| 4 (Likely) | 4 | 8 | 12 | 16 | 20 |
| 5 (Certain) | 5 | 10 | 15 | 20 | 25 |

| COLOUR DEFINITIONS |  |
| --- | --- |
| GREEN | Keep risk register to demonstrate awareness of risk |
| AMBER | Efforts should be made to reduce the risk and risk reduction measures implemented where possible within a defined time period. Further assessment may be necessary to establish more precisely the likelihood of harm as a basis for determining the need for improved control measures. |
| RED | Very detailed risk assessment needs to be carried out and risks must be reduced. |

| TERMINOLGY DEFINITIONS |  |  |  |
| --- | --- | --- | --- |
| Probability |  | Severity |  |
| 1 (Remote) | Extremely unlikely to occur in equipment life | 1 (Negligible) | Very unlikely to sustain any injury |
| 2 (Unlikely) | Unlikely to occur in equipment life | 2 (Minor) | Minor injury (first aid not required), repair not required |
| 3 (Possible) | Could occur | 3 (Moderate) | Minor injury first aid required, some damage that will need repair |
| 4 (Likely) | Likely to occur | 4 (Severe) | Significant injury possible that requires hospital treatment |
| 5 (Certain) | Will happen in equipment life | 5 (Significant) | Major injury or death |

| Wheelchair Mounting |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ref | Hazard | Pre-management Risk Score |  |  | Notes and Management (delete / edit as needed) Solutions included in this column are suggestions only and should be used as a guide | Post-management Risk Score |  |  |
|  |  | S | P | T |  | S | P | T |
| 1 | Chair tipping due to changes in weight distribution. |  |  |  | Advise client to remove device when mobilising in environments identified as presenting a tipping risk (e.g. uneven terrain, unfamiliar terrain, slopes or hills) |  |  |  |
| 2 | Pressure injury / bruising sustained due to unintentional movements or resting posture resulting in contact with the device or mounting system. |  |  |  | Adjust the position of the device AND / OR Pad the mounting or device to reduce pressure risk |  |  |  |
| 3 | Injury sustained if the means of propelling the chair (either manually or under power) is obstructed by the mounting system or device in any position. |  |  |  | Identify other means of wheelchair propulsion when device is mounted. AND / OR Provide written instruction only to propel the chair when the device is in a specified position that reduces obstruction |  |  |  |
| 4 | Injury sustained if the user propels the chair (either manually or under power) and their vision is obscured by the mounting system or device in any position. |  |  |  | Provide written and verbal instructions to be aware and take account of blind spots when driving. AND / OR Advise upon need for supervision of driving, or removal of device, in high-risk situations e.g. busy or unfamiliar environments |  |  |  |
| 5 | Chair being involved in a collision due to the mounting system extending beyond the dimensions of the wheelchair to the front, rear or sides, affecting safe clearance. |  |  |  | If changing overall dimensions cannot be avoided, provide written and verbal instructions to be aware of additional dimensions of the chair |  |  |  |
| 6 | Injury sustained or device damaged due to chair tipping when device and mount is left in the swung away position / side folded position. |  |  |  | Consider alternative provision of a static mounting pole. OR Provide written and verbal instructions to always remove the device if leaving mount in swung-away position |  |  |  |
| 7 | Injury sustained due to trapping body parts between the mounting or device and the wheelchair, or in moving mechanisms of the mount. |  |  |  | Minimum gap of (2cm) between components in all positions. OR Label of hazard area clearly. OR identification in documentation or at handover. OR Provide written and verbal warnings regarding the danger areas for entrapment and to instruct carers in safe fitting and removal of the device and mount. |  |  |  |
| 8 | Injury sustained if device and mount are left in position during transit and move when subjected to dynamic forces of driving |  |  |  | Provide written and verbal advice to always remove and safely store both the device and mount before transit. AND Provide an alternative, low tech method of communication for use during transit OR Complete a risk assessment for transport |  |  |  |
| 9 | Injury sustained or device damaged due to nuts, bolts or other fittings become loose over time and mounting fails |  |  |  | Provide verbal and written instructions to regularly check the mounting components and what action should be taken. |  |  |  |
| 10 | Injury sustained or device damaged if the device is incorrectly mounted and / or not locked in position |  |  |  | Ensure client and/or carer are competent in removing the device and mount. AND Provide written instructions with photographs and diagrams. |  |  |  |
| … |  |  |  |  |  |  |  |  |

| Desk / Table / Floor Mount |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ref | Hazard | Pre-management Risk Score |  |  | Notes and Management (delete / edit as needed) Solutions included in this column are suggestions only and should be used as a guide | Post-management Risk Score |  |  |
|  |  | S | P | T |  | S | P | T |
| A | Injury sustained or device damaged due to the device being incorrectly mounted and / or not locked in position. |  |  |  | Provide training and ensure that client/carers/others are competent in fitting/removing the device. AND Consider advising to not remove the device if situation dictates. AND / OR Consider use of different mounting system for easier fitting. |  |  |  |
| B | Pressure injury / bruising sustained due to unintentional movements or resting posture resulting in contact with the device or mounting system. |  |  |  | Adjust the position of the device. AND / OR Pad the mounting or device to reduce pressure risk. |  |  |  |
| C | Injury sustained or device damaged due to the floor mount tipping over when device is left in the swung away position |  |  |  | Provide verbal and written instructions to remove the device before folding the mount. AND / OR Ensure that the swing-away function does not leave the device outside the footprint of the floor mount. |  |  |  |
| D | Injury sustained or device damaged due to the weight of the device causing the mounting to fail, particularly if the mount is being used outside of the manufacturer’s recommendation(s) for weight and positioning. |  |  |  | Ensure that the mount is appropriately weight-rated for the device being mounted. AND Only mount within the footprint of the mount. AND Provide verbal and written warnings advising client/carers/others to not adjust the position outside of the footprint. |  |  |  |
| E | Injury sustained by others tripping over / colliding with the mount |  |  |  | Provide verbal and written warnings to highlight potential trip-hazard. AND Apply hazard tape to floorstand base / legs. AND / OR Identify location for storage when not in use. AND / OR Consider additional risk management in contexts where there are lots of other people or clients with diagnoses such as visual impairment or behavioural. |  |  |  |
| F | Injury sustained by persons supporting the user due to moving or lifting the device or mount improperly |  |  |  | Minimise need to move the mount and the distances over which it needs to be moved – consider the provision of additional mounts if practical. AND / OR Ensure that those supporting the user have relevant and up-to-date manual handling training in all environments where the mount will be used. AND / OR Provide training specific to the mount used. |  |  |  |
| G | Injury sustained or device damaged due to the mount obstructing moving part(s) of a profiling bed or chair. |  |  |  | Provide verbal and written warnings highlighting the risk of obstruction and what positions to avoid when using the device/mount. |  |  |  |
| H | Damage to mount caused by inappropriate storage when not in use |  |  |  | Provide verbal and written instructions on appropriate storage. AND / OR Ensure that an appropriate and agreed designated storage area is identified in each location where the mount will be used. |  |  |  |
| … |  |  |  |  |  |  |  |  |

# Section 5 – Handover Information

| HANDOVER DETAILS |  |  |
| --- | --- | --- |
| Handover Date | Handover Location | Handover By |

|  | CONTACT INFORMATION |  |  |  |
| --- | --- | --- | --- | --- |
|  | Name | Organisation | Phone | Email |
| Device Provider |  |  |  |  |
| Named Receiving Person |  |  |  |  |
| Handover Contact / First Line Support |  |  |  |  |

HANDOVER NOTES

# APPENDIX 1 – Wheelchair Static Stability Test Results

| Wheelbase / Wheelchair Reference (from main MAT-DOC): |  |
| --- | --- |
| Front Track Width: |  |
| Rear Track Width: |  |

| SUMMARY |  |  | LOAD CELL DATA (WHERE USED) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Ref | Test 1 |  |  |  | Left (kg) | Right (kg) |
| Configuration | Before install |  | Front up | Front |  |  |
| Notes |  |  | Level | Front |  |  |
|  |  |  |  | Rear |  |  |
|  |  |  | Front down | Rear |  |  |
| RESULTS |  |  |  |  |  |  |
| Uphill |  | Downhill | Sideways (L down) |  | Sideways (R down) |  |

| SUMMARY |  |  | LOAD CELL DATA (WHERE USED) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Ref | Test 2 |  |  |  | Left (kg) | Right (kg) |
| Configuration | After install |  | Front up | Front |  |  |
| Notes |  |  | Level | Front |  |  |
|  |  |  |  | Rear |  |  |
|  |  |  | Front down | Rear |  |  |
| RESULTS |  |  |  |  |  |  |
| Uphill |  | Downhill | Sideways (L down) |  | Sideways (R down) |  |

# APPENDIX 2 – Driving Assessment

| Wheelbase / Wheelchair Reference (from main MAT-DOC): |  |
| --- | --- |

| BACKGROUND INFORMATION |  |
| --- | --- |
| Brief description of environment and conditions in which the assessment is taking place |  |
| Brief description of user’s driving abilities prior to provision of mounting |  |
| Any special considerations, relating to driving, identified during pre-provision process? |  |

| PRE-ASSESSMENT CHECKS |  |
| --- | --- |
| Can the client see over the top of the mounted equipment? |  Yes  No |
| Can the client see underneath the mounted equipment? |  Yes  No |
| Is there any visible damage to the environment which might indicate prior accidents or collisions? |  Yes  No |
| Is there any visible damage to the wheelchair which might indicate prior accidents or collisions? |  Yes  No |

| OBSERVATIONS |  |  |  |
| --- | --- | --- | --- |
| Were any issues observed when… | Without device | With device | Comments (Including details of how items were assessed, any risks identified during assessment and how these risks will be managed) |
| …negotiating narrow openings? |  Yes  No |  Yes  No |  |
| …manoeuvring in areas with limited space? |  Yes  No |  Yes  No |  |
| …noticing and avoiding obstacles? |  Yes  No |  Yes  No |  |
| …noticing and avoiding other people in the environment? |  Yes  No |  Yes  No |  |
| …approaching and stopping next to desk / table without knocking into it? |  Yes  No |  Yes  No |  |
| …following instructions when driving (e.g. “Stop”) without being distracted? |  Yes  No |  Yes  No |  |
| Other (please specify) |  Yes  No |  Yes  No |  |